## R2C-Asbestos
Response to Comments - ELV1M3 Quality System for Asbestos Testing
Disclaimer: The NELAC Institute (TNI) accepts no liability for the content of any comment on a standard.
Any views or opinions on a standard are solely those of the commenter and do not necessarily reflect those of TNI.
Notice of Intent
| Date of CSDEC Proposal/Approval |  |  | Date of Website Posting |  |  | Comment Closure Date |  |  |  | Suggested Langauge (if Provided)
| 9/8/17 & 9/27/17 |  |  | 9/28/17 |  |  | 10/28/17
Comment number | Vote (Persuasive/non-persuasive) | Justification | Editorial (Y/N) | Section/ clause | Comment | Committee action | Date addressed | Committee comment | Date Submitted | Date & Method of Submitter Notification
No comments received upon publication of NOI and/or ANSI announcement
Summary and Outline of Proposed Changes
|  |  |  | Date of Website Posting |  |  | Comment Closure Date
|  |  |  | 2019-07-30 |  |  | 2019-08-30
No comments received upon publication of Summary and Outline of Proposed Changes
Presentation of Proposed Draft Standard for TNI Comment; Forum on Environmental Accreditation, Newport BeacH, CA, Thursday 2/6/2020
Comment number | Vote (Persuasive/non-persuasive) | Justification | Editorial (Y/N) | Section/ clause | Comment | Committee action | Date addressed | Committee comment | Date Submitted | Date & Method of Submitter Notification | Suggested Langauge (if Provided)
1 | Non-persuasive |  | No | 3.1 | Add definition of "equivalent" | None | 2020-04-15 | Committee felt definition not necessary beyond common understanding and text of language provided with standard materials | 2020-02-06 | 5/21/20 by e-mail from R. Wyeth
2 | Non-persuasive | Term familiar to field of asbestos resting | No | 3.1 | clarification of terms "Molecular % and Optical Properties" | None | 2020-04-15 | Terms clearly understood in asbestos testing industry and is appropriated referenced in the module | 2020-02-06 | 5/21/20 by e-mail from R. Wyeth
3 | Non-persuasive | Related documents section is not presented in Standard Volumes and/or Modules | No | NA | SOP 1-104 should be added to module as will a section on related documents and a module history | None | 2020-05-20 | Committee on 4/15/20 believed comment non-persuasive but module history will be added upon completion of the Standard Module. | 2020-02-06 | 5/21/20 by e-mail from R. Wyeth
4 | Non-Persuasive | Refernce is correct | No | 7.2.1.4 | Was reference citation in this section correct? | None | 2020-04-15 | Citation was verified as correct | 2020-02-06 | 5/21/20 by e-mail to all meeting participants
5 | Persuasive | Reference is correct but will be expanded | No | 7.2.3.4 | Was reference citation in this section correct? | Expand section in draft ELV1M3 | 2020-04-15 | Citation was verified as correct but committee for completeness added an additional reference [EPA 600/R-93/116(1993)] | 2020-02-06 | 5/21/20 by e-mail to all meeting participants
6 | Persuasive | As certificates are required is reference necessary | No | 7.4.7.c | Questioned applicability of citing ANSI N42.22-1995. | Citation was removed here and in 7.2.4.3.c | 2020-05-20 | Committee agreed citation not required as certificates must be retained by laboratory | 2020-02-06 | 5/21/20 by e-mail to all meeting participants
Summary of Changes and Justification; Publication of DS and opening of Comment Period (TNI and Non TNI members
|  | Date of Website Posting |  |  | Comment Clousure Date |  |  | ANSI submission of BSR-8 & announcement in Standards Action | ANSI comment period closure Date
|  | 2020-12-22 |  |  | 2021-03-24 |  |  | Verified 6/7/2021, announced 6/18/2021 | 8/2/2021; no comments received
Comment number | Vote (Persuasive/non-persuasive) | Justification | Editorial (Y/N) | Section/ clause | Comment | Committee action | Date addressed | Committee comment | Date Submitted | Date & Method of Submitter Notification | Submitter Comment/Suggestion
7 | Non-Persuasive | Asbestos analysis utilizes only standard methods, any lab developed methods are generally proposed to become standard methods. | No | 5 | The use of the term "non-standard methods" has always prompted questions on what is a non-standard method, what is a standard method (generally accepted as such), and when does a modification to a standard method make this method non-standard. | none | 2021-04-21 | Terms clearly understood in asbestos testing industry | 2021-03-08 | By e-mail 9/20/2021 | Last Sentence: Change to: "Laboratory-developed test methods must comply with the requirements in Volume 1, Module 2, Section 5.4.5.
8 | Persuasive | Clarification on phrasing is appropriate | No | 6.1.c) | There is confusion on what the Demonstration of Capability requirements are for the laboratory to be accredited, and what the requirements are for laboratory analysts in order to determine competence. Sections 6.1(a) and 6.1(c) appear to pertain to laboratory individuals whereas 6.1(b) and 6.1(d) could apply to either individuals or to the laboratory as a whole. Many EPA test methods require laboratory initial demonstrations of capability to be available and on file regardless of how long ago they were performed (i.e., records retention beyond the 5-year minimum) unless the laboratory repeated the Initial Demonstration of Capability. Records available for only an ongoing demonstration of capability for the laboratory will not be acceptable. | Revise draft ELV1M3 | 2021-04-21 | Although all DOC's in the asbestos industry are individual not lab the committee does feel the clarification is justified. | 2021-03-08 | By e-mail 9/20/2021 | First sentence: Add to the end, as follows: "In cases where an individual has prepared and/or analyzed samples using a method that has been in use by the laboratory for at least one (1) year prior to applying for accreditation, and there have been no significant changes in instrument type or method, the on-going DOC shall be acceptable as an initial DOC for that individual."
9 | Persuasive | Valid | No | 6.3.2(a) | Provision of a method blank or a sample with non-detect concentrations of Asbestos is unacceptable for a blind sample to be used in a demonstration of capability. | Revise draft ELV1M3 | 2021-04-21 | It is common for real word samples to result in no asbestos detected and these samples should not be used to demonstrate capability | 2021-03-08 | By e-mail 9/20/2021 | Add to (a): "acceptable performance of a blind sample (single blind to the analyst and containing Asbestos) or successful analysis of a blind sample on a similar method using the same technology (e.g., EPA Methods 100.1 and 100.2);"
10 | Persuasive | Valid | No | 7.2.1.1(d) | No frequency is specified for the beam dose calibrations. | Revise draft ELV1M3 | 2021-04-21 | Missed in draft | 2021-03-08 | By e-mail 9/20/2021 | Add this as the last sentence to the Section (consistent with requirements in NIST NVLAP and in NY-ELAP 198.4): "This calibration shall be performed quarterly."
11 | Persuasive | Valid | Yes | 7.2.1.1(f) | Acronyms should be spelled out upon first usage. Also, the equipment to be calibrated appears to be incorrectly named. | Revise draft ELV1M3 | 2021-04-21 | Missed in draft | 2021-03-08 | By e-mail 9/20/2021 | Revise first complete sentence as follows: "In accordance with AHERA requirements (Asbestos Hazard Emergency Response Act) (40 CFR Part 763, Subpart E, App. A), the low temperature asher shall be calibrated quarterly for 5-10% mass loss by …"
12 | Persuasive | Valid | Yes | 7.2.1.2 Note | The acronym MFL shows up for the first time here, but it is not spelled out or defined until Section 7.2.4.1. Is this OK? | Revise draft ELV1M3 | 2021-05-19 | Missed in draft | 2021-03-08 | By e-mail 9/20/2021 | If not OK, then add to the NOTE: "See Section 7.2.4.1 for definition of MFL."
13 | Persuasive | Valid | Yes | 7.2.1.2(c) | The complete reference for "Turner and Steel" is provided later in Section 7.2.4.2(a). | Revise draft ELV1M3 | 2021-05-19 |  | 2021-03-08 | By e-mail 9/20/2021 | Add to the parentheses at the end of the 2nd sentence: "(NISTIR 5351, see Section 7.2.4.2(a))."
14 | Non-Persuasive | repeating method requirements is redundant | No | 7.2.1.3 | The analytical sensitivity depends on more than the fraction of the filter sampled and dilution factor. | None | 2021-05-19 | Committee believes the standard need not repeat method requirements and that labs should refer to the method. While commenter is not wrong the committee felt the suggestion was overly perscriptive and redundant relative to the method. | 2021-03-08 | By e-mail 9/20/2021 | Revise the last sentence as follows: "This value will depend on the amount of water sample filtered, the effective filter area of the sampling filter, the number of TEM grid openings counted, the area of the TEM grid openings, and the dilution factor (if applicable)."
15 | Non-Persuasive | repeating method requirements is redundant | No | 7.2.1.4(a) | The published EPA Method Errata is not incorporated into this Standard for the EPA 100.2 calculation formula. | None | 2021-05-19 | Committee believes the standard need not repeat method requirements and that labs should refer to the method. While commenter is not wrong the committee felt the suggestion was overly perscriptive and redundant relative to the method. | 2021-03-08 | By e-mail 9/20/2021 | Add to the clause as follows: "The concentration of asbestos in a given sample shall be calculated in accordance with EPA/600/R-94/134, Method 100.2, Section 12.1 (with the GOA measurement units corrected to mm2) or Method 100.1, Section 7.
16 | Persuasive | valid | No | 7.2.2.1 | Should the standard presentation be more clearly presented, with incomplete details added? | Revise draft ELV1M3 | 2021-05-19 | will add language as proposed | 2021-03-08 | By e-mail 9/20/2021 | Express this section as follows: "All calibrations shall be performed in accordance with Section 7.2.1.1, with the exception of: a) Magnification used for fiber counting shall be 15,000 to 20,000x, and b) Calibration of the camera constant shall be performed at least monthly."
17 | Persuasive | Will add to clarify | Yes | 7.2.2.2.(b) | Does the reader know what, or why, NIST SRM 1876b is required for this section? | Revise draft ELV1M3 | 2021-05-19 | Anyone perfrming Asbestos Testing should know! But for sake of new/less experienced analysts could add. Committee decided not necessary but will add for clarification | 2021-03-08 | By e-mail 9/20/2021 | For information, add in the first sentence: "… shall obtain mean analytical results on NIST SRM 1876b (Chrysotile Asbestos) or equivalent so that trimmed mean values …"
18 | Persuasive | Will add to clarify | No | 7.2.2.2(e) | What are the acceptance criteria for the Replicate analysis in (i) and for the Verified Analyses in (iii)? | Revise draft ELV1M3 | 2021-05-19 | no need to add language as requirement to evaluate perfromance covered in 7.2.1.2(a). But will add for clarification Also for consistency harmonize all elements of the 7.2.2.2. | 2021-03-08 | By e-mail 9/20/2021 | Add to (i): "Replicate. A second, independent analysis shall be performed with results evaluated for acceptance in accordance with Section 7.2.1.2(a);" Add to (iii): "Verified Analyses. A second, independent analysis shall be performed on the same grids and grid openings with results evaluated for acceptance in accordance with Section 7.2.1.2(c)."
19 | Persuasive | valid | No | 7.2.2.2(f) | Acceptance criteria have not been specified in this Standard. | Revise draft ELV1M3 | 2021-05-19 | will add the language | 2021-03-08 | By e-mail 9/20/2021 | Add a second sentence as follows: "Criteria for acceptability shall be established and documented by the laboratory."
20 | Non-Persuasive | repeating method requirements is redundant | No | 7.2.4.1(a)(ii) | For EPA Method 100.2, Section 1.6 specifies an analytical sensitivity for samples at 0.2 MFL; however, Section 9.1 of this same method requires that the process blanks have < 0.01 MFL asbestos greater than 10 um in length. To accomplish this sensitivity, the laboratory must count more TEM grid squares and/or filter larger volumes of blank fiber-free water. 200-mesh TEM grids are 0.080 mm on a side, so the grid opening area (GOA) is 0.0064 mm2. From EPA 100.1 (Chatfield, et. al.), the effective filter area (efa) of a 25-mm diameter filter is 1990 mm2, and the efa of a 47-mm diameter filter is 11340 mm2. That means that the product of (# of grid openings counted) x (sample volume filtered in mL) must be at least 33100 and 177200, respectively. | None | 2021-05-19 | procedure is referenced in method. Committee believes standard is adequate but may want to revisit for additional clarity and lab useability. Ultimiately viewed non-persuasive. | 2021-03-08 | By e-mail 9/20/2021 | Suggests re-write of section, has not but will contribute proposed language
21 | Persuasive | valid | No | 7.2.4.2(a)(i) | Some laboratories have not received 100 drinking water Asbestos samples in a long time. Should there be a minimum timeframe frequency also be specified for analyzing 1 of these Asbestos preparations? Also, for analyzing all 6 Asbestos preparations? | Revise draft ELV1M3 | 2021-05-19 | will insert the language | 2021-03-08 | By e-mail 9/20/2021 | Add as a last sentence: "At a minimum, all 6 Asbestos preparations shall be analyzed at least once every two years."
22 | Persuasive | valid | No | 7.2.4.2(b)(i) | Once the asbestos preparations are made, the Standard does not have any clear requirements to analyze them and to evaluate the results. | Revise draft ELV1M3 | 2021-05-19 | Suggested..."Referece filter preparations shall be made from all six asbestos types in accordance with …. | 2021-03-08 | By e-mail 9/20/2021 | Add to the sentence in the clause as follows: "Filter preparations shall be made, analyzed, and the results evaluated from all six (6) asbestos types in accordance with Section 7.2.4.2(a)(i).
23 | Persuasive | Valid | Yes | 7.2.4.3(a) | It may not be appropriate for a "QC program" to be the subject of this sentence as written. | Revise draft ELV1M3 | 2021-05-19 | Add clarity | 2021-03-08 | By e-mail 9/20/2021 | Add to the sentence as follows: "The laboratory in its QC program shall establish and maintain provisions for asbestos standards."
24 | Persuasive | Valid | No | 7.4.2.(a) | Requirements for Intra-Analyst Precision are specified, but the acceptance criteria are not specified | Revise draft ELV1M3 | 2021-05-19 | Add clarity | 2021-03-08 | By e-mail 9/20/2021 | Add a last sentence to this section: "The laboratory shall establish and document acceptance criteria consistent with the requirements in Section 7.4.4.2."
25 | Persuasive | Valid | No | 7.4.5(a) & 7.4.5(b) | Requirements are specified for Friable Materials and for Non-Friable Materials, but in both cases acceptance criteria are not specified. | Revise draft ELV1M3 | 2021-09-15 | This section refers to Blanks, add language to state if contamination detected, no analysis can proceed and the source of contamiation to be investigated and resolved if possible. | 2021-03-08 | By e-mail 9/20/2021 | Add a new Section 7.4.5(c) as follows: "For both Section 7.4.5(a) and Section 7.4.5(b), the laboratory shall establish and document acceptance criteria consistent with the Analytical Sensitivity requirements in Section 7.4.3."
26 | Persuasive | Valid | Yes | 7.4.6(a) | The second sentence is written in the passive voice, so it is not evident who is supposed to be implementing the requirements in this Standard. | Revise draft ELV1M3 | 2021-05-19 | Add clarity | 2021-03-08 | By e-mail 9/20/2021 | Revise the second sentence as follows: "The laboratory shall analyze, out of one hundred (100) samples, at least one (1) standard or reference sample that has been routinely resubmitted to determine the analyst's precision and accuracy, as developed by the laboratory in accordance with quality control requirements."
27 | Persuasive | Valid | Yes | 7.4.6(b) | The sentence is written in the passive voice, so it is not evident who is supposed to be implementing the requirements in this Standard. | Revise draft ELV1M3 | 2021-05-19 | Add clarity | 2021-03-08 | By e-mail 9/20/2021 | Revise the sentence as follows: "Non-Friable Materials. The laboratory shall analyze at least one (1) out of one hundred (100) samples with a verified quantitative standard that has routinely been resubmitted to determine analyst precision and accuracy."
28 | Persuasive | Valid | Yes | 7.4.7(A) | It may not be appropriate for a "QC program" to be the subject of this sentence as written. | Revise draft ELV1M3 | 2021-05-19 | Add clarity | 2021-03-08 | By e-mail 9/20/2021 | Add to the sentence as follows: "The laboratory in its QC program shall establish and maintain provisions for asbestos standards."
29 | Persuasive | Valid | Yes | 7.5.3 | The sentence is written in the declarative voice; thus, it is not evident to whom the requirement refers. | Revise draft ELV1M3 | 2021-05-19 | Add clarity | 2021-03-08 | By e-mail 9/20/2021 | Rewrite the sentence as follows: "The laboratory shall refer and conform to the specific method of analysis for additional requirements."
30 | Non-Persusaive | Will search | Yes | Global | search for the first use of an acronym, assure that it is spelled out then, and then don’t repeat that process (see PLM, for example, which is spelled out in 1.0, and then again in 7.4.1.1) | Will revise if necessary | 2021-06-16 | Missed in Draft | 2021-03-23 | By e-mail 9/20/2021
31 | Persuasive | Agree with commenter | Yes | 5 | I believe you should strike the ‘and’ in the last line. Non-standard methods must comply with the requirements in Volume 1, Module 2, Section 5.4.5. As an aside, this reference will change as Module 2 is revised. It MIGHT be better to refer to something such as ‘…the requirements for Method Validation found in TNI EL Volume 1 Module 2.’ This naming allows for the requirements to move, but not get lost due to a reference change. | Revise draft ELV1M3 | 2021-06-16 | go with 1st option; remove "and" | 2021-03-23 | By e-mail 9/20/2021
32 | Persuasive | Agree with commenter | Yes | 6.1 a) | – this section refers to a Section 1.6.2. I believe you mean Section 6.2 of your revised Module. | Revise draft ELV1M3 | 2021-06-16 | will correct | 2021-03-23 | By e-mail 9/20/2021
33 | Non-Persuasive | Standard operating procedure for Asbestos testing lab | No | 6.1 d) | is the 5 year retention period from the time that the DOC was performed, or 5 years after the analyst has stopped performing the analysis, or from some other starting point? I suggest that it be 5 years after the analyst has stopped performing the analysis. | none | 2021-06-16 | std ops for labs, no change | 2021-03-23 | By e-mail 9/20/2021
34 | Non-Persuasive | understood by AT labs | no | 6.2.1 c) | – is ‘class of analyte’ the equivalent of the type of asbestos, such as chrysotile? Should that term be defined or is there another term that can be used? | None | 2021-06-16 | class of analyte recognized in AT | 2021-03-23 | By e-mail 9/20/2021
35 | Non-Persuasive | NA | No | 6.3.1 & 6.3.2 d) | seem to support that Ongoing DOC may be achieved by reviewing an analyst’s routine work, in addition to the requirements in 6.3.2 a) through c). If that is the case, I wholeheartedly support that language. | None | 2021-06-16 | thanks | 2021-03-23 | By e-mail 9/20/2021
36 | Non-Persuasive | No clarification necessary | No | 7.2.1.1. | includes the phrase ‘may have to be increased’. Does the Committee have a recommendation on how often calibrations have to be performed following non-routine maintenance or unacceptable calibration | None | 2021-06-16 | Believe the std is sufficently clear | 2021-03-23 | By e-mail 9/20/2021
37 | Persuasive | Valid | No | 7.2.1.1 b) | requires calibrations be performed weekly. If a calibration performed Monday, and another performed the following Friday is acceptable, then I have no issue with this wording. If that isn’t the intent, I would suggest stating ‘every 7 days’. | Revise draft ELV1M3 | 2021-06-16 | need to clarify; every 7 days (+/- 2 days) | 2021-03-23 | By e-mail 9/20/2021
38 | Persuasive | Valid | No | 7.2.1.1 e) iii) | refers to calibration semiannually. This term “typically” means every 6 months. I suggest stating every 6 months rather than semiannually. In the same sentence, if the detector geometry is not altered, could calibration be less frequent than every 6 months? | Revise draft ELV1M3 | 2021-06-16 | need to clarify; every 6 months (+/- 1 month) | 2021-03-23 | By e-mail 9/20/2021
39 | Non-Persuasive | not Valid | No | 7.2.2.1 b) | requires calibrations be performed monthly. Similar to my comment above, if a calibration performed on the first day of the month, and another performed on the last day of the next month is acceptable, then I have no issue with this wording. If that isn’t the intent, I would suggest stating ‘every 30 days’. | Revise draft ELV1M3 | 2021-06-16 | need to clarify; every 30 days (+/- 5 days) | 2021-03-23 | By e-mail 9/20/2021 | there is no 7.2.2.1.b)
40 | Persuasive | Agree with commenter | Yes | 7.2.2.2 b) | refers to 7.2.1.2 c below. It precedes this section, so it is above. I suggest that the reference is sufficient without referring to above or below. | Revise draft ELV1M3 | 2021-06-16 | change to "abov"e | 2021-03-23 | By e-mail 9/20/2021
41 | Persuasive | Agree with commenter | Yes | 7.2.3.2 a), b), and c) | the use of ‘at least’ and ‘at a minimum’ is unnecessary, as the entire Standard is the minimum that must be done. (see also any other instances of these or ‘at minimum’ through the Standard). | Revise draft ELV1M3 | 2021-06-16 | will remove | 2021-03-23 | By e-mail 9/20/2021
42 | Non-persuasive | non-Valid | No | 7.2.3.3 | doesn’t contain anything that is either a requirement or assessable. Is that the intent? If so, it should be changed to a NOTE | None | 2021-06-16 | method dependent and not requiring definition for bulk analysis | 2021-08-17 | By e-mail 9/20/2021 | Suggested leanguage "if any asbestos is detected, the source of comtamination shall be traced and the problem resolved to prevent reoccuranve"
43 | Persuasive | Agree with commenter | Yes | 7.2.3.4 | ‘with’ appears to be missing between ‘accordance’ and ‘the’ | Revise draft ELV1M3 | 2021-06-16 | change | 2021-03-23 | By e-mail 9/20/2021
44 | Persuasive | Agree with commenter | Yes | 7.2.4.1 a) ii | the correct abbreviation for milliliter is mL, not ml. It appears that the section refers to l (“ell”) liter, rather than 1 liter in the 4th line | Revise draft ELV1M3 | 2021-06-16 | change | 2021-03-23 | By e-mail 9/20/2021
45 | Non-Persuasive | contained in method | No | 7.2.4.1 b) i | refers to polycarbonate blank filters and blank MCE filters. Is it the intent that differing types of filters have different blank criteria? That isn’t clear from this section | None | 2021-07-21 | yes they are different but this is included within the method…clear to all knowledgeable of abs analysis | 2021-03-23 | By e-mail 9/20/2021
46 | Persuasive | Agree with commenter | No | 7.2.4.2 a) ii | refers to analysis by each analyst. Section 7.2.4.2 b)ii does not refer to analysis by each analyst. Is that intentional? | Revise draft ELV1M3 | 2021-07-21 | will add analysis by each analyst | 2021-03-23 | By e-mail 9/20/2021
47 | Persuasive |  | Yes | 7.3.1.3 | the formatting at the end needs to be corrected. The 2 in mm2 needs to be a superscript. I would advise to be consistent in the use of + 2 µm or plus or minus 2 micron (my preference is the former). | Revise draft ELV1M3 | 2021-07-21 | editorial will fix | 2021-03-23 | By e-mail 9/20/2021
48 | Persuasive | Clarification of terms | No | 7.3.2 b) | references ‘well-behaved’ samples. I might suggest using the term ‘typical’ or ‘routine’ instead. Does this requirement apply to each analyst who performs air counts on a given day, or each analyst who MAY perform air counts on a given day? In the former, there may only be 1 analyst. In the latter, it would require all to participate daily. | Revise draft ELV1M3 | 2021-07-21 | change behaved to "characterized…requirement is for each analyst who performs on a given day | 2021-03-23 | By e-mail 9/20/2021
49 | Persuasive | Agree with commenter | Yes | 7.3.3 | is it an ideal range on the counting filter, or must it be 100 to 1300 fibers/mm2? An ideal range indicates that it could be outside of that range. Section 7.3.4 seems to indicate that results outside of that range must be qualified, in which case I’m not sure that the ‘ideal range’ statement is necessary | Revise draft ELV1M3 | 2021-07-21 | remove the work ideal | 2021-03-23 | By e-mail 9/20/2021
50 | Persuasive | Valid | No | 7.4.1.2 | by requiring semiannual calibration, do you mean twice per year, or once every 6 months? If either are ok, semiannual is fine. | Revise draft ELV1M3 | 2021-07-21 | need to clarify; every 6 months (+/- 1 month) | 2021-03-23 | By e-mail 9/20/2021
Presentation of Revised(2nd) Proposed Draft Standard for TNI Comment; Forum on Environmental Accreditation, San Antonio, TX January 2022
Comment number | Vote (Persuasive/non-persuasive) | Justification | Editorial (Y/N) | Section/ clause | Comment | Committee action | Date addressed | Committee comment | Date Submitted | Date & Method of Submitter Notification | Suggested Langauge (if Provided)
51 | Non-persuasive | Valid | Yes | Throughout | Monthly was explicitly stated earlier in the Module. Should it be done so again? | Revise draft ELV1M3 | 2022-02-16 | Include where appropriate | 2022-01-17 | NA - Public Comment
52 | Non-persuasive | Not Valid | Yes | Throughout | there are places where 'at least' is the correct usage, so as to not make a requirement where one isn't necessary. | No change | 2022-02-16 | Each notation is correct in DS | 2022-01-17 | NA - Public Comment
5300% | Non-persuasive | Not valid | Yes | Throughout | this seems like a conflicting requirement. Is it every 6 months if less frequent is every 12 months because that is 'next use'? | Revise draft ELV1M3 | 2022-02-16 | define 6 months as + or - 1 month where necessary | 2022-01-17 | NA - Public Comment